## Slide 1: Metrology Automation Lunch & Learn CLS & MET/CAL® Drivers
Michael L. Schwartz

## Slide 2
OOP Object Oriented Programming approach applied to a non-OOP language
        See an early version of the Metrology Taxonomy the NCSLI 141 Committee is currently working on
        Get some FREE sample code
Learning Objectives

## Slide 3
Veteran owned business, ex-military metrologist
Core business is software
  Metrology consulting
  Efficiencies through automation
  Turn-key system
Focus on finding the right solution for the customer
  Created Metrology.NET® in 2015
  World’s largest MET/CAL® procedure library
  PS-Cal® power sensor calibration solution - Now Tegam’s Product
  C# web-based asset management system
Acquired Cal Lab Magazine in 2011
About Cal Lab Solutions

## Slide 4: Business is about Efficiency?
Better
Cheaper
Faster
That doesn’t mean less accurate

## Slide 5
Cal Lab Solutions writes custom software for companies around the world. We could not effectively write and support our procedures if every procedure had the lab standards hard coded in the MET/CAL® procedure.
We needed a way to write automation that could be quickly customized to any calibration lab’s existing standards.
We needed a way to create automated calibration procedures: Better – Highest quality Cheaper – Lower support costs Faster – Needs to be < 10x the manual calibration time
Problem Statement

## Slide 6
Overview
Main Procedures - UUT Specific
  ID the UUT Organize execution order Organize print order
  Call Sub Test Points
Sub Test Points - UUT Specific
  List all test points Evaluate Pass / Fail Call Sub Test Routines
Sub Test Routines - UUT Specific
  Set up the UUT Send Generic Commands to the Driver Call Sub Config
Sub Config – Station / Customer Specific
  Give Connection Messages Call Specific Drivers
CLSD Cal Lab Solutions Driver – Standard Specific
  Translates Generic Commands to Instrument Commands Calculates Uncertainty

## Slide 7
Free Example Procedure Keysight DSO-X 4000A
Main Procedures - UUT Specific
Sub Test Points - UUT Specific
Sub Test Routines - UUT Specific
Sub Config – Station Specific
CLSD Cal Lab Solutions Drivers
https://store.callabsolutions.com/index.php?route=product/product&product_id=55
Fluke 5520A - Version
Fluke 5820A – Version Fluke 9500 w/ 4 Heads – Version
Fluke 5720 / 33250 / PSG / E4419 / E9304A – Version
}

## Slide 8
#======================= * End to End Sequencer * ============================
2.047 JMPL Start_sequencer
2.048 EVAL Copyright 2011 Cal Lab Solutions, LLC. All rights reserved.
3.001 LABEL Start_sequencer
# Channel 1 Test
3.002 JMPL DCVerticalGain_1
3.003 LABEL DCVerticalGain_1_done
3.004 JMPL DualCursor_1
3.005 LABEL DualCursor_1_done
#============================================================================
9.001 LABEL DCVerticalGain_1
9.002 RSLT =
9.003 HEAD { * DC VERTICAL GAIN * }
9.004 HEAD * DC VERTICAL GAIN { Channel 1 } *
9.005 IF Find(S[23],"Skip_DCVerticalGain",1)==0
9.006 MATH S[24]="DCVerticalGain,1"
9.007 CALL Agilent DSOX-4000 Sub Test Points
9.008 ELSE
9.009 RSLT = Skipped
9.010 ENDIF
9.011 JMPL TestSelectionMode (L[30]==1)
9.012 JMPL DCVerticalGain_1_done
9.013 EVAL Copyright 2011 Cal Lab Solutions, LLC. All rights reserved.
Code Dive Main Procedure

## Slide 9
1.001 MATH M[20]=Fld(S[24],2,",")
1.004 MATH L[1]=Find(S[24],"DCVerticalGain",1)==1
1.005 JMPL DCVerticalGain (L[1]==1)
#============================================================================
9.001 MATH S[25]="Check Connections"
9.002 LABEL DCVerticalGain
9.003 MATH S[26]="ReConfig"
# Perform Any Pre-Calibration Steps
9.004 CALL Agilent DSOX-4000 Sub Test Routines-Init
#-----------------------------
9.005 MATH M[1]=35.000*1
9.006 MATH M[2]=5*1
9.007 TSET UUT_Res = .001
9.008 CALL Agilent DSOX-4000 Sub Test Routines-Conf
9.009 MATH L[9]=Fld(S[31],2,"Unc=")
9.010 ACC 35.000V L9U
9.011 IF 1==0
9.012 TARGET -m
9.013 ENDIF
9.014 CALL Agilent DSOX-4000 Sub Test Routines-Meas
9.015 MATH MEM=Fld(S[31],2,"Value=")
9.016 MEMCX V 0.8U 5V/Div
Code Dive Test Points

## Slide 10
1.001 MATH M[20]=Fld(S[24],2,",")
1.004 MATH L[1]=Find(S[24],"DCVerticalGain",1)==1
1.005 JMPL DCVerticalGain (L[1]==1)
#============================================================================
9.001 MATH S[25]="Check Connections"
9.002 LABEL DCVerticalGain
9.003 MATH S[26]="ReConfig"
# Perform Any Pre-Calibration Steps
9.004 CALL Agilent DSOX-4000 Sub Test Routines-Init
#-----------------------------
9.005 MATH M[1]=35.000*1
9.006 MATH M[2]=5*1
9.007 TSET UUT_Res = .001
9.008 CALL Agilent DSOX-4000 Sub Test Routines-Conf
9.009 MATH L[9]=Fld(S[31],2,"Unc=")
9.010 ACC 35.000V L9U
9.011 IF 1==0
9.012 TARGET -m
9.013 ENDIF
9.014 CALL Agilent DSOX-4000 Sub Test Routines-Meas
9.015 MATH MEM=Fld(S[31],2,"Value=")
9.016 MEMCX V 0.8U 5V/Div
Code Dive Test Points

## Slide 11
3.011 MATH S[30]="Check Connection"
3.012 CALL Agilent DSOX-4000 Sub Config
# Configure to Testing
3.013 IF Find(S[26],"DCVerticalGain"&M[20],1)==0
# Reset & Configure Measurement Device
3.014 MATH S[30]="Reset Channel= " & M[20]
3.015 CALL Agilent DSOX-4000 Sub Tools
# Reset & Configure Source
3.016 MATH S[30]="Source.Volts.DC Reset"
3.017 CALL Agilent DSOX-4000 Sub Config
# Setup Measurement Device
3.018 MATH S[30]="Write:TIM:SCAL 10e-3"
3.019 CALL Agilent DSOX-4000A Communication
3.030 MATH S[26]="DCVerticalGain"&M[20]
3.031 ENDIF
Code Dive Test Routine

## Slide 12
# Setup Source
3.036 MATH S[30]="Source.Volts.DC "
3.037 MATH S[30]=S[30]&" Volts= "&M[1]
3.038 CALL Agilent DSOX-4000 Sub Config
3.039 MATH L[9]=Fld(S[31],2,"Unc=")
# Measure ==============================
3.046 LABEL DCVerticalGain_Meas
# Measure the Value
3.047 MATH S[30]="Query:MEAS:RES?"
3.048 CALL Agilent DSOX-4000A Communication
# Calculate Required Values
3.049 MATH MEM=Fld(MEM2,6,",")
3.050 TSET U3 = [MEM]
#Return Results
3.051 MATH MEM=Fld(MEM2,5,",")
3.052 MATH S[31]="Value= "&MEM&" Unc="&L[9]
3.053 END
Code Dive Test Routine

## Slide 13
8.001 LABEL DCVerticalGain
8.002 JMPL DCVerticalGain_Conn Find(S[30],"Connect",1)>0
8.003 JMPL Source.Volts.DC Find(S[30],"Source",1)>0
8.004 DISP Error Calling Sub
8.005 END
#========================================
# Required Specifications
# Volts DC 14mV to 35V
#============================================================================
2.001 LABEL Source.Volts.DC
2.002 MATH S[30]=S[30]&" Channel= "&M[20]
2.003 CALL CLSD-Source.Volts.DC (9500 w/ Head(s))
2.004 END
Code Dive Sub Config

## Slide 14
#========================================
8.006 LABEL DCVerticalGain_Conn
8.007 IF Fld(S[8],1,",")==4
8.008 IF Find(S[25],"9500.4Ch<--->UUT.4Ch",1)==0
8.009 DISP Connect the Fluke 9500 to the UUT as Follows:
8.009 DISP [32] Channel 1 <---> Channel 1
8.009 DISP [32] Channel 2 <---> Channel 2
8.009 DISP [32] Channel 3 <---> Channel 3
8.009 DISP [32] Channel 4 <---> Channel 4
8.009 DISP
8.009 DISP * Do not use any terminations.
8.010 MATH S[25]="9500.4Ch<--->UUT.4Ch"
8.011 ENDIF
8.012 ELSE
8.013 IF Find(S[25],"9500.2Ch<--->UUT.2Ch",1)==0
8.014 DISP Connect the Fluke 9500 to the UUT as Follows:
8.014 DISP [32] Channel 1 <---> Channel 1
8.014 DISP [32] Channel 2 <---> Channel 2
8.014 DISP
8.014 DISP * Do not use any terminations.
8.015 MATH S[25]="9500.2Ch<--->UUT.2Ch"
8.016 ENDIF
8.017 ENDIF
Code Dive Sub Config

## Slide 15
#========================================
8.001 LABEL DCVerticalGain
8.002 JMPL DCVerticalGain_Conn Find(S[30],"Connect",1)>0
8.003 JMPL Source.Volts.DC Find(S[30],"Source",1)>0
8.004 DISP Error Calling Sub
8.005 END
#========================================
8.006 LABEL DCVerticalGain_Conn
8.007 IF Find(S[25],"5520SC<--->UUT.Channel"&M[20],1)==0
8.008 DISP Connect the 5520 Scope Output to the
8.008 DISP UUT Channel [M20] Input.
8.009 MATH S[25]= "5520SC<--->UUT.Channel"&M[20]
8.010 ENDIF
8.011 END
#============================================================================
# Required Specifications
# Volts DC 14mV to 35V
#============================================================================
2.001 LABEL Source.Volts.DC
2.002 CALL CLSD-Source.Volts.DC (5520A Scope)
2.003 END
Code Dive Sub Config

## Slide 16
2.026 LABEL Source.Volts.DC
# Get Channel
2.027 IF (Findi(S[30],"Channel=",1)>0)
2.028 MATH L[1]=fld(S[30],2,"Channel=")
2.029 ELSE
2.030 MATH L[1]=1
2.031 ENDIF
# Get Volts
2.032 IF (Findi(S[30],"Volts=",1)>0)
2.033 MATH L[2]=fld(S[30],2,"Volts=")
2.034 ELSE
2.035 DISP Missing Voltage: CLS Driver Source.DC (4 Head 9500)
2.036 JMPL Error
2.037 ENDIF
# Volts Into 50 Ohms
2.047 IF (L[3]==50)
2.048 IF (L[2]<=-5.56)||(L[2]>=5.56)
2.049 DISP Volts out of range in call to CLS Driver
2.049 DISP Source.DC (4 Head 9500)
2.049 DISP [L2] Volts (+/- 5.56 Limit into 50 Ohms)
2.050 JMPL Error
2.051 ENDIF
2.052 ENDIF
Code Dive Driver

## Slide 17
# Set Output
#-----------
2.062 IEEE [@9500]ROUTE:SIGN CH[L1]
2.063 IEEE [@9500]FUNC DC
2.064 IEEE [@9500]ROUT:SIGNAL:IMPEDANCE [L3]
2.065 IF L[2]==0
2.066 IEEE [@9500]PAR:DC:GROUND ON
2.067 ELSE
2.068 IEEE [@9500]PAR:DC:GROUND OFF
2.069 IEEE [@9500]VOLT [L2]
2.070 ENDIF
2.071 IEEE [@9500]OUTPUT ON
#Error Check the 9500
2.072 IEEE [@9500]SYST:ERR?[i$]
2.073 IF MEM2!=0
2.074 MATH MEM2 = SUB(MEM2,1,68)
2.075 DISP 9500 Reported the Following Error
2.075 DISP [MEM2]
2.076 ENDIF
Code Dive Driver

## Slide 18
# Return Level, Uncertainties
#--------------------------
2.077 IF L[3]==50
2.078 MATH L[15]=Accv("Datron 9500","Volts L",L[2])
2.079 ELSEIF L[3]==1e6
2.080 MATH L[15]=Accv("Datron 9500","Volts",L[2])
2.081 ENDIF
2.082 MATH S[31]= "Value="&L[2] &" Unc="&L[15]
# Expanded Uncertainties
#---------------------------------
# Standard Resolution
2.087 IF L[2]==0
2.088 MATH L[31] = .01e-3
2.089 ELSE
2.090 MATH L[30]=10e-6
2.091 WHILE Abs(L[2])>L[30]
2.092 MATH L[30]=L[30]*10
2.093 ENDW
2.094 MATH L[30]=L[30]/1e5
2.095 MATH L[31] = L[30]
2.096 ENDIF
2.097 MATH L[31]=L[31]/2/Sqrt(3)
2.098 TSET U7 = [L31]
# Standard Traceability (Assuming 4 to 1 or Better)
2.099 MATH L[31]=L[15]*.25
2.100 TSET U5 = [L31]
Code Dive Driver

## Slide 19
1.001 MATH M[20]=Fld(S[24],2,",")
1.004 MATH L[1]=Find(S[24],"DCVerticalGain",1)==1
1.005 JMPL DCVerticalGain (L[1]==1)
#============================================================================
9.001 MATH S[25]="Check Connections"
9.002 LABEL DCVerticalGain
9.003 MATH S[26]="ReConfig"
# Perform Any Pre-Calibration Steps
9.004 CALL Agilent DSOX-4000 Sub Test Routines-Init
#-----------------------------
9.005 MATH M[1]=35.000*1
9.006 MATH M[2]=5*1
9.007 TSET UUT_Res = .001
9.008 CALL Agilent DSOX-4000 Sub Test Routines-Conf
9.009 MATH L[9]=Fld(S[31],2,"Unc=")
9.010 ACC 35.000V L9U
9.011 IF 1==0
9.012 TARGET -m
9.013 CALL Agilent DSOX-4000 Sub Test Routines-Meas
9.014 ENDIF
9.015 MATH MEM=Fld(S[31],2,"Value=")
9.016 MEMCX V 0.8U 5V/Div
Code Dive Test Points

## Slide 20
OOP Object Oriented Programming approach applied to a non-OOP language
        See an early version of the Metrology Taxonomy the NCSLI 141 Committee is currently working on
        Get some FREE sample code
Learning Objectives

## Slide 21: Questions? / Comments
Michael L. Schwartz
Cal Lab Solutions
mschwartz@callabsolutions.com
https://store.callabsolutions.com/index.php?route=product/product&product_id=55